Breathalyzer Firmware Instructions and Configuration Guide

Project: Breathalyze-BG2.0  |  Target board: Adafruit Feather ESP32 V2

# 1. Firmware overview

This document explains how the current firmware operates, how to build and use it, which source-level variables are intended to be adjusted, how QR and BLE payloads are formed, how sensor sampling works, and how the firmware state machine moves from power-up to deep sleep.

·        Platform and framework: ESP32 on the Adafruit Feather ESP32 V2 using PlatformIO and the Arduino framework.

·        Display and input: 170×320 ST7789 TFT, one physical button, and a CST816 touch controller used mainly on the glucose-entry screen.

·        Sensing model: five analog sensors are sampled continuously and stored in a ring buffer; baseline values and two breath-test result sets are extracted from that history.

·        Output paths: the firmware can send data over BLE, render a QR code for the final URL payload, and save payload-only test history in LittleFS flash storage.

·        Primary control file: src/BreathalyzerApp.cpp contains the user interface, state machine, sensor capture logic, QR generation, BLE orchestration, and power-off behavior.

# 2. Project structure and module responsibilities

| File | Purpose |
| --- | --- |
| platformio.ini | Build environment, target board, filesystem selection (LittleFS), partition table, and external library dependencies. |
| src/main.cpp | Hardware pin map, watchdog setup, creation of the HardwareConfig object, and top-level setup()/loop(). |
| src/app_config.h | User-adjustable timing, display, battery, and payload constants. |
| src/BreathalyzerApp.cpp / .h | Main application implementation: UI, state machine, sensor sampling, QR generation, BLE payload routing, glucose entry, and deep sleep. |
| src/BleUart.cpp / .h | BLE server, client-mode detection, RX parsing, text notification chunking, and 45-value binary packet generation. |
| src/ButtonInput.cpp / .h | Debounced physical-button events such as tap, press-down, and hold. |
| src/TouchCst816.cpp / .h | Touch-controller probe, polling readout, coordinate scaling, inversion, and fault backoff. |
| src/DataStore.cpp / .h | Legacy NVS blob save/load plus payload-history ring buffer in LittleFS. |

# 3. How to build, flash, and use the firmware

## 3.1 Build and flash workflow

1.      Open the PlatformIO project. The uploaded firmware is a PlatformIO project configured for the adafruit_feather_esp32_v2 environment.

2.      Install the required libraries. The project's dependencies in the Platform.ini file are:

[env:adafruit_feather_esp32_v2]

platform = espressif32

board = adafruit_feather_esp32_v2

framework = arduino

monitor_speed = 115200

board_build.filesystem = littlefs

board_build.partitions = no_ota.csv

lib_deps =

ricmoo/QRCode @ ^0.0.1

adafruit/Adafruit GFX Library@^1.12.4

adafruit/Adafruit ST7735 and ST7789 Library@^1.11.0

h2zero/NimBLE-Arduino @ ^2.1.0

3.      Connect the board and upload. Use a normal PlatformIO Upload action to compile and flash the firmware.

4.      Open the serial monitor at 115200 baud. Serial is useful for BLE debugging, QR URL inspection, and the DUMP/ERASE maintenance commands.

## 3.2 Runtime operation for the device user

5.      Power on. Hold the physical button for approximately 900 ms to boot from deep sleep or from a fresh power-on reset.

6.      Warmup. The display shows a warmup page and progress bar while the heater is enabled for 10 seconds.

7.      Test 1 ready screen. Tap the button to start the first blow. On the first test, the button press also marks the baseline-capture reference point.

8.      Blow for 6 seconds. The firmware shows a countdown and stores the time-aligned 4 s, 5 s, and 6 s result samples for the test.

9.      Intermission. A 10-second intermission screen appears before the second test.

10.   Test 2. Tap the button again and complete the second 6-second blow.

11.   Glucose entry / BLE stage. After test 2 the heater turns off, BLE advertising starts, and the user can either enter BG manually on the touchscreen or let the mobile app send BG and timestamp data.

12.   QR stage. If manual BG entry is used, the firmware builds the final URL payload and renders it as a QR code for scanning.

13.   Power off. Hold the button for about 3 seconds to force deep sleep. The firmware can also auto-sleep on timeout or go to a short Done screen when a BLE app response completes the workflow.

·        Serial maintenance commands while device is on: Typing DUMP followed by Enter prints saved payload history; typing ERASE wipes the stored history and legacy NVS namespace.

·        Auto-off note: The current AUTO_OFF_AFTER_MS default is (120,000 ms ≈2 mins). It is active in the Ready, Glucose, and ShowQr phases.

·        Touch usage note: Touch is used only on the glucose-entry screen; the physical button drives most other state changes.

# 4. Firmware flowchart description

The following text flow is a description of the firmware flowchart implemented in BreathalyzerAppImpl::appBegin() and BreathalyzerAppImpl::appUpdate().

[Deep sleep / Power off]
     	|
 Hold button for ON_HOLD_TO_BOOT_MS
     	v
   appBegin() + displayInit()
     	|
     	v
  	Warmup (heater on)
     	|
 WARMUP_DURATION_MS elapsed
     	v
  	Ready - Test 1
     	|
  button tap captures candidate baseline point
     	v
   	Blow - Test 1
     	|
  BLOW_WINDOW_MS elapsed
     	v
  	Intermission
     	|
 INTERTEST_DELAY_MS elapsed
     	v
  	Ready - Test 2
     	|
     	v
   	Blow - Test 2
     	|
  save payload to LittleFS history
     	v
    Glucose Entry + BLE advertising
   	/                	\
  	/                  	\
 BLE BG/timestamp     	Manual BG entry
 received              	by touch/button
 	|                      	|
 	v                      	v
   Done screen           	Show QR screen
 	|                      	|
   3 s sleep         	hold-to-off / auto-off
 	\__________________________/
              	|
              	v
         	Deep sleep

·        Global transition #1: Holding the button longer than OFF_HOLD_TO_SLEEP_MS from normal runtime screens forces goToDeepSleep(true).

·        Global transition #2: If the displayed battery percentage falls to 1% or lower while the heater is on for roughly 1.5 seconds, the firmware shows LOW BATT and enters deep sleep.

·        Global transition #3: After the tests are complete, a valid BLE BG response can interrupt Glucose or ShowQr and move the firmware to the Done screen immediately.

·        Global transition #4: On startup, if the wake/reset path does not satisfy the hold-to-boot check, the firmware returns to deep sleep instead of continuing.

# 5. State machine details

| State | What the user sees | Main actions in the state | Exit condition |
| --- | --- | --- | --- |
| Warmup | WARMUP page with progress bar | Heater on, progress percentage updated, button hold-to-off still available. | When the elapsed warmup time reaches WARMUP_DURATION_MS; then a double beep and transition to Ready. |
| Ready | TEST 1 or TEST 2 page with start prompt | Waits for stable button release, allows a tap to start the blow, and for Test 1 captures the candidate press time/values used for baseline extraction. | Button tap enters Blow. Auto-off also applies here. |
| Blow | BLOW NOW!! with countdown | Continuously samples sensors, shows remaining seconds, captures 4/5/6 second result values at the end of the blow. | At BLOW_WINDOW_MS expiration: Test 1 goes to Intermission; Test 2 saves payload history and goes to Glucose. |
| Intermission | REST / INTERMISSION countdown | Wait state between tests while sensor sampling continues in the background. | At INTERTEST_DELAY_MS expiration, the firmware beeps once and returns to Ready for Test 2. |
| Glucose | BG# manual entry screen | BLE advertising starts if not already started, user can edit three digits with touch arrows, and physical button or touch OK confirms BG. | Manual confirmation enters ShowQr; BLE response after tests complete can instead send the firmware to Done. |
| ShowQr | QR screen | Builds and renders the QR payload once, continues BLE activity, and can redraw the QR on a button tap. | Auto-off or hold-to-off enters deep sleep; BLE response after tests complete can still redirect to Done. |
| Done | ALL DONE / GOODBYE | Short terminal state after a BLE app response provides BG and timestamp. | After about 3 seconds, enters deep sleep. |

## Important state-machine guard variables

·        _offArmed / ARM_RELEASE_STABLE_MS: Prevent the same boot press from being interpreted as an immediate power-off request.

·        _readyArmed / READY_ARM_RELEASE_MS: Prevent a held or bouncing button from skipping through the Ready state right after transition from Intermission or startup.

·        _testsComplete, _savedThisRun, _savedSeq: Track when the second test has finished and which saved flash record should be updated later with BG/timestamp.

·        _doneSleepAtMs and _autoOffDeadlineMs: Manage post-completion and inactivity-based deep sleep.

·        _currentTest: Selects whether the firmware is operating on Test 1 or Test 2.

# 6. Sensor polling, ring buffer, and result extraction

The sensor path is centered in sampleSensorsToRingIfDue(), ringGetAtTime(), captureBaselineOnceFromCandidate(), captureResultsAt456(), and buildSudirValues().

·        Polling rate: All five analog sensor channels are read every SAMPLE_PERIOD_MS (20 ms by default). The ADC reading is converted to volts with raw * 3.3 / 4095.

·        Ring buffer: The firmware keeps a ring of SensorSample objects in _ring with RING_SIZE = 220 entries, which is about 4.4 s of history at the default sample period.

·        Baseline capture: On the first test, the candidate press time is marked when the user presses the button on the Ready screen. Once the blow starts, the firmware extracts three baseline time points: 2 s before the press, 1 s before the press, and the sensor values at the press itself.

·        Result capture: For each blow, three result snapshots are extracted at approximately 4 s, 5 s, and 6 s from _blowStartMs. If a target sample is missing, the most recent sample is used as a fallback.

·        Data model: Baseline values are stored in _baseVals[5][3]. Result values are stored in _resVals[2][5][3] for two tests and three time points per sensor.

·        BLE binary packing: buildSudirValues() flattens the 45 values (15 baseline + 30 results) into uint16 numbers after multiplying volts by 1000 and clipping at 65535.

·        API/QR scaling: The URL builder uses the same scale-by-1000 rule and sends each sample as an integer field value.

## Why ringGetAtTime() works with a 6-second blow even though the ring stores about 4.4 seconds

The result capture happens at the end of the blow. At that moment, the 4-second sample is only about 2 seconds old, the 5-second sample is about 1 second old, and the 6-second sample is current. All three therefore fit comfortably inside the 4.4-second ring buffer.

# 7. QR-code generation details

The QR path is implemented by chooseQrVersionForLength(), buildResultUrlCommon(), generateQr(), and drawQrOnBlack().

·        QR content source: QR data comes from buildResultUrlQR(), which calls buildResultUrlCommon(..., true) so the QR payload includes the BASE_URL plus query parameters.

·        Payload-only variants: buildResultPayload() and buildResultUrlBT() both currently call buildResultUrlCommon(..., false), so flash storage and BLE text notifications use payload-only strings without the BASE_URL prefix.

·        Field order: The URL begins with d=<API_DEVICE_ID> and p=<API_PARAM_P>. If available, bg=<three-digit glucose> is added next. Sensor groups follow, then ts=<timestamp> if a timestamp is present.

·        Sensor group names: Sensor groups are mapped as S1→a, S2→b, S3→c, S4→e, and S5→f. There is intentionally no group letter d because d is reserved for device ID in the URL.

·        Per-sensor indexing: For each sensor group, fields 1..3 are baseline values, 4..6 are Test 1 values, and 7..9 are Test 2 values.

·        QR version selection: chooseQrVersionForLength() walks a fixed low-error-correction capacity table and picks the smallest QR version that can hold the current payload length.

·        Error correction level: The qrcode library is called with ECC_LOW.

·        Rendering style: drawQrOnBlack() creates a centered QR with a 4-module quiet zone, scales it to the current display, and draws white modules on a black square. This is effectively an inverted QR style.

·        Header text: The QR screen prints SCAN QR CODE near the top of the display.

·        Current flag behavior: USE_SUDIR_QR is present in app_config.h, but in the current code both branches of enterShowQr() call buildResultUrlQR(); the flag is therefore acting as a placeholder rather than selecting a different QR builder today.

## Example payload layout

BASE_URL + "d=18&p=BG&bg=123&a1=...&a2=...&...&a9=...&b1=...&...&f9=...&ts=<unix>"

# 8. BLE firmware details

BLE behavior is handled by the BleUart class. The main application calls _ble.begin("BGBT_FEATHER_V2") when it reaches the Glucose or ShowQr phase. begin() is idempotent, so repeated calls are safe.

·        Device name: BGBT_FEATHER_V2.

·        Current active UUID set: The code currently uses the 6E400021 / 6E400022 / 6E400023 service/RX/TX UUID triplet. A second Sudir UUID set is present in comments but is not active.

·        Server behavior: The firmware creates one NimBLE server, one service, one notify/read TX characteristic, and one write/write-no-response RX characteristic. Advertising restarts on connect and disconnect.

·        TX text path: sendTextPayload() uses notifyChunked() with a 160-character chunk size so long strings are emitted as several notifications.

·        TX binary path: sendBinaryPacket() builds a 92-byte packet: 2 bytes of device ID followed by 45 little-endian uint16 values.

·        Client modes: The code distinguishes Unknown, Logan, and Sudir client behavior.

·        Logan ASCII handshake: If the RX side receives printable text containing READY, the client is classified as Logan and the firmware raises a ready-to-send flag. On the next maybeSendBlePayload() pass, the text payload is notified with a newline.

·        Logan BG response: Printable text containing BG=### is parsed as a valid app response. Optional timestamp keys TS=, TIME=, TIMESTAMP=, &ts=, &time=, or &timestamp= are also recognized.

·        Sudir response forms: A single byte 0xAA is treated as a Sudir ACK. A 6-byte write is interpreted as little-endian BG (2 bytes) plus timestamp (4 bytes).

·        Unknown-client grace period: If a client stays connected longer than UNKNOWN_TO_SUDIR_GRACE_MS (1200 ms) without Logan-style activity, shouldBroadcastBinary() begins treating it as a binary-broadcast candidate.

·        Runtime usage in app: During Glucose and ShowQr, maybeSendBlePayload() decides whether to send Logan text payloads or Sudir-style binary packets based on the detected client mode. But currently, in Logan's App for both Android and iOS, all payloads will be Logan-style, and Sudir-style payloads are just a placeholder that can be removed if they cause a conflict.

## Important BLE implementation note

The current code uses payload-only text strings for BLE text notifications because buildResultUrlBT() calls buildResultUrlCommon(..., false). If the mobile app expects the full BASE_URL, this function is the one to change.

# 9. Flash storage and persistence details

·        Legacy NVS save/load: DataStore::save() and DataStore::load() still support a SavedRun blob in the Preferences namespace "bgbt".

·        Current main-history path: The active runtime flow uses appendQuery() and updateQuery() instead of the blob save/load path.

·        Filesystem: History is stored in LittleFS at /bgbt_history.bin. LittleFS.begin(true) is used so the filesystem can be formatted if mounting fails.

·        History format: A small header tracks magic, version, capacity, count, start index, record size, and next sequence number. Each record stores a sequence number and one payload string.

·        Retention rule: MAX_SAVED_TESTS is 1000. When the history is full, the firmware overwrites the oldest payload in ring-buffer fashion.

·        Stored content: Only the payload string is stored; the BASE_URL is intentionally not written to flash.

·        Update behavior: After Test 2, the payload is saved immediately. Later, when manual glucose entry or BLE timestamp data becomes available, updateSavedQueryWithLatestResult() rewrites the same record with the final BG/timestamp values.

·        Maintenance commands: DUMP prints stored payloads in oldest-to-newest order. ERASE clears the Preferences namespace and deletes the LittleFS history file.

# 10. Touch, button, power, and battery behavior

## 10.1 Button input

·        Debounce model: ButtonInput defaults to a 30 ms debounce period and treats taps shorter than MIN_TAP_MS (40 ms) as invalid.

·        Events provided: stable press-down edge, tapped release, hold event, and current down/not-down state.

·        Hold event behavior: wasHeld() is one-shot per press; once the hold event fires, it will not refire until the button is released and pressed again.

## 10.2 Touch input

·        Touch mode: TouchCst816::begin() defaults to polling mode instead of interrupt mode because the configured INT pin is GPIO37, which does not support internal pullups.

·        Probe and setup: The firmware retries the CST816 I2C probe for about one second and disables the controller auto-sleep register so touch stays responsive.

·        Scaling and orientation: RAW_MAX_X = 170, RAW_MAX_Y = 320, SWAP_XY = false, INVERT_X = true, and INVERT_Y = true are used to map raw touch data to screen coordinates.

·        Failure handling: If I2C reads fail repeatedly, the code backs off and eventually disables touch for the rest of the boot session to stop Wire error spam.

## 10.3 Battery and sleep

·        Battery source: Battery voltage is measured through the Feather battery monitor pin and converted back to pack voltage by doubling the averaged ADC millivolt reading.

·        Display model: The current code intentionally updates the displayed percentage while the heater is on and freezes it when the heater is off, so the number reflects loaded device voltage rather than battery open-circuit voltage.

·        Current battery mapping: BATT_V_FULL = 4.00 V maps to 100% and BATT_V_EMPTY = 3.20 V maps to 0% using a linear calculation. Mapping this battery voltage range to a non-Linear voltage discharge graph still needs to be done for the most accurate time% to battery% approximation.

·        Low battery cutoff: If the displayed battery percentage is 1% or below while the heater is on for about 1.5 seconds, the firmware shows LOW BATT and enters deep sleep.

·        Wake source: Deep sleep uses ext0 wake on the button pin, with polarity chosen from the button-active-low configuration.

# 11. Changeable variables and what they do

This section focuses on the constants and configuration points that are practical to modify. Not every local variable is listed; the emphasis is on values that affect timing, hardware mapping, payload format, storage, or protocol compatibility.

## 11.1 Recommended user-adjustable constants in src/app_config.h

| Variable | Default | What it controls / what happens if changed |
| --- | --- | --- |
| SERIAL_BAUD | 115200 | Serial monitor speed for debug messages and maintenance commands. |
| ON_HOLD_TO_BOOT_MS | 900 ms | How long the power button must be held to boot from sleep/power-on. Lower values make startup easier; higher values reduce accidental wakes. |
| OFF_HOLD_TO_SLEEP_MS | 3000 ms | How long the button must be held during runtime to force deep sleep. |
| WARMUP_DURATION_MS | 120000 ms | Total heater warmup time before the first test is allowed. |
| INTERTEST_DELAY_MS | 10000 ms | Rest time between Test 1 and Test 2. |
| AUTO_OFF_AFTER_MS | 120000 ms | Inactivity timeout for Ready, Glucose, and ShowQr. |
| BLOW_WINDOW_MS | 6000 ms | How long the user must blow in each test. Also drives the countdown and 4/5/6-second capture points. |
| SAMPLE_PERIOD_MS | 20 ms | Sensor polling period. Reducing it increases time resolution and CPU usage. Increasing it lowers resolution. |
| PREBASE_NEED_MS | 2200 ms | Minimum required history before baseline extraction is considered valid. |
| ARM_RELEASE_STABLE_MS | 120 ms | Button release stabilization time before hold-to-off becomes armed after boot. |
| READY_ARM_RELEASE_MS | 120 ms | Button release stabilization time before the Ready state starts accepting a tap-to-start event. |
| TFT_W / TFT_H | 170 / 320 | Logical display size passed to the ST7789 driver and touch scaling assumptions. |
| TFT_ROTATION | 0 | Display rotation used by the ST7789 driver. |
| TFT_SPI_HZ | 40000000 | Display SPI bus speed. Lower it if the display becomes unstable. |
| DEVICE_ID_STR | "D1" | Declared string identifier. It is currently not used by the main payload builders. |
| DEVICE_ID_NUM | 0x0020 | Two-byte device ID placed at the front of the 92-byte binary BLE packet. |
| API_DEVICE_ID | 18 | Device identifier placed in the URL/QR field d=. |
| API_PARAM_P | "BG" | Value placed in the URL/QR field p=. |
| USE_SUDIR_QR | true | Currently a placeholder flag; both branches still call buildResultUrlQR() in the present code. |
| BASE_URL | Supabase function URL | Base URL added only for QR generation in the current firmware. |

## 11.2 Hardware mapping constants in src/main.cpp

| Constant | Default | When to change it |
| --- | --- | --- |
| PIN_BUTTON | 32 | Change if the wake/power button is rewired to a different RTC-capable GPIO. |
| PIN_BUZZER | 8 | Change if the buzzer moves to a different output. |
| PIN_HEATEREN | 7 | Change if the heater-enable transistor or switch uses another pin. |
| TFT_SCLK / TFT_MOSI / TFT_MISO | 5 / 19 / 21 | Change if the TFT is moved off the current Feather SPI wiring. |
| TFT_CS / TFT_DC / PIN_TFT_BL / PIN_TFT_RST / PIN_TFT_PWR | 14 / 33 / 27 / 4 / 15 | Display chip-select, data/command, backlight, reset, and display-power control pins. |
| PIN_BATT_MON | 35 | Battery monitor ADC pin. Change only if using a different board or divider node. |
| PIN_I2C_SCL / PIN_I2C_SDA | 20 / 22 | Touch I2C bus routing. |
| PIN_TOUCH_INT | 37 | Touch interrupt pin. If moved to a pullup-capable GPIO, interrupt mode becomes more practical. |
| SENSOR_PINS | {26, 25, 34, 39, 36} | Five analog sensor-channel pins. |
| BUTTON_ACTIVE_LOW | false | Switch to true if the button is wired active-low instead of active-high. |
| USE_INTERNAL_PULLS | false | Enable only if the button needs internal pullups/pulldowns instead of an external resistor network. |
| WDT_TIMEOUT_S | 15 | Task watchdog timeout used in main.cpp. |

## 11.3 Practical tuning constants in src/BreathalyzerApp.cpp

| Variable / constant | Default | Effect |
| --- | --- | --- |
| RING_SIZE | 220 | Amount of sensor history kept in RAM. Increase if you need more look-back time; decrease if RAM is tight. |
| Battery ADC averaging N | 8 samples | Higher values smooth the battery display more but take longer to read. |
| BATT_V_EMPTY | 3.20 V | Voltage that maps to 0% battery in the current loaded-voltage model. |
| BATT_V_FULL | 4.00 V | Voltage that maps to 100% battery in the current loaded-voltage model. |
| Low-battery hold time | 1500 ms | How long battery must remain at <=1% before forced sleep. |
| TOP_H / BOTTOM_OVERLAY_H | 36 / 34 | Top-header area and bottom live-overlay reservation on the TFT. |
| BATT_W / BATT_H / BATT_NUB_W | 36 / 10 / 5 | Visual size of the battery icon. |
| Warmup easing profile | 0–90% in first 75%, eased 90–100% in last 25% | Change warmupPercent() if you want a different progress-bar feel. |
| Glucose UI sizes | boxW=44, boxH=64, gap=10, OK button 120×44 | Adjust if the glucose-entry screen layout needs to be resized. |
| QR quiet zone | 4 modules | Change drawQrOnBlack(4) if you want a larger or smaller quiet border. |
| QR version table | QR_BYTE_CAPACITY_LOW[] | Update only if changing QR error-correction assumptions or library behavior. |
| Live overlay rate | 250 ms | Refresh period for drawLiveSensorOverlay() when that debugging overlay is enabled. |

## 11.4 BLE protocol constants in src/BleUart.cpp

| Variable | Default | Meaning |
| --- | --- | --- |
| SERVICE_UUID | 6E400021-...CCA9E | Active BLE service UUID used today. |
| CHAR_RX_UUID | 6E400022-...CCA9E | Write characteristic UUID used by the phone app. |
| CHAR_TX_UUID | 6E400023-...CCA9E | Notify/read characteristic UUID used by the firmware. |
| TOTAL_VALUES | 45 | Number of uint16 sensor values in the Sudir-style binary packet. |
| PACKET_SIZE | 92 bytes | Derived from 2-byte device ID plus 45 values × 2 bytes each. |
| ACK_VALUE | 0xAA | Single-byte Sudir ACK token. |
| UNKNOWN_TO_SUDIR_GRACE_MS | 1200 ms | How long an unknown client can remain connected before binary broadcasting starts. |
| Text chunk size | 160 chars | Chunk size for notifyChunked() when sending long text payloads. |
| ESP power level | ESP_PWR_LVL_P9 | BLE radio power. Lower if power draw or RF aggressiveness must be reduced. |

## 11.5 Storage, touch, and button constants

| File / variable | Default | Meaning |
| --- | --- | --- |
| DataStore::MAX_SAVED_TESTS | 1000 | Maximum payload history entries retained in LittleFS before overwriting the oldest entry. |
| DataStore::MAX_PAYLOAD_LEN | 480 | Maximum payload string length stored per history record. |
| Touch RAW_MAX_X / RAW_MAX_Y | 170 / 320 | Used to scale raw CST816 coordinates into screen coordinates. |
| Touch SWAP_XY / INVERT_X / INVERT_Y | false / true / true | Coordinate orientation and inversion settings. |
| Touch INT_ACTIVE_LOW | true | Expected interrupt polarity if interrupt mode is used. |
| Touch i2cHz parameter | 100000 | Touch I2C clock passed into begin(); useful if bus speed must change. |
| Button debounceMs | 30 ms default | Debounce interval used by ButtonInput::begin(). |
| ButtonInput::MIN_TAP_MS | 40 ms | Minimum press length treated as a valid tap. |

# 12. Known implementation notes and team-facing cautions

These are artifacts left in the code from previous iterations that used multiple endpoints for BLE and QR code that shouldn't affect what the code does, just make reading what the code does more confusing:

·        Header declaration note: BreathalyzerApp.h declares buildResultUrlQRSudir(), but there is no active implementation in the current source tree.

·        QR flag note: USE_SUDIR_QR is not currently selecting a different QR payload path; both branches call the same builder.

·        BLE text note: The BLE text payload currently excludes BASE_URL. If the mobile app expects a full URL, change buildResultUrlBT().

Device changes to consider:

·        Touch pin note: GPIO37 is input-only and has no internal pullup, which is why the firmware defaults to polling rather than touch-interrupt mode.

·        Battery-display note: The displayed percent is currently a device-under-load estimate, not a true chemistry-based state-of-charge gauge. Needs range to be mapped to a voltage discharge graph of the battery being used to account for non-linearity. Didn't have enough time to get this discharge graph of the battery, but for use, the battery percent still functions as an accurate indicator of battery at 100%, 50%, and 0%, with the largest discrepancy between the linear and non-linear discharge being at 99%-90% and 10%-1%.

·        Storage note: The LittleFS history path is designed for payload-only archival. The older SavedRun blob code still exists, but is not the main path used after Test 2.

·        UI note: Several countdown and percentage displays are now drawn by updating only the number field to reduce flicker; if screen layout changes, those cursor positions may need to be retuned.

·        Compatibility note: Changing BLE UUIDs, URL field names, API_DEVICE_ID, packet size, or sensor-group naming may also require mobile-app and backend changes.

# 13. Summary

The firmware is organized around a single application state machine that warms the sensors, captures two timed breath tests, lets the user or app provide glucose data, and then publishes the result through BLE, QR, and flash storage. Most day-to-day configuration changes live in app_config.h and main.cpp. More invasive changes—such as protocol changes, sensor-mapping changes, or alternate QR/BLE payload formats—are centered in BreathalyzerApp.cpp and BleUart.cpp.